Annex 2 : Measurement uncertainty

A2.1	Introduction

Measurement uncertainty arises from a number of causes and is often classified into random and systematic uncertainty. The former results from the combined effects of a number of small independent random variables in the measuring system, while the latter can be the result of factors such as instrument calibration, measurement conditions (e.g. temperature), etc. The combined effect is to produce a range of values within which the true value of a measured result actually lies. Thus the value of the quantity to be measured can be expressed as a probability distribution and the usual mathematical techniques applied to the analysis.

An appreciation of these factors is of value in understanding the general problem and in ensuring that the smallest uncertainties are achieved in a particular test apparatus. For testing of GSM mobile stations a system simulator is used and the various circuits of this equipment and the associated measuring equipment contribute to the total uncertainty in any measurement. These elements have been identified and the estimations are set out in Annex 5 at a confidence level of 95%.

Many of the uncertainties arise in the RF area where cables, connectors and switches with differing attenuation, VSWR, mismatch loss, etc, occur between the measuring point and the MS antenna terminal. These RF paths are systematically changed during testing and this causes further uncertainty. Also the input impedance of a mobile under test will not necessarily be known in advance which gives rise to possible additional mismatch loss, and the impedance at 12.5 GHz will almost certainly not be the same as at 900 MHz.

Given that the overall uncertainty is composed of several factors and that the random combination of a number of distributions tends towards a Gaussian distribution advantage can be taken of the relatively simple mathematical analysis that this affords.

The RF tests set out in GSM 11.10, sections 12, 13, 14 and 15 through 22, audio tests in section 30.1 and the stimulus setting levels in other tests, are the tests that are most susceptible to measurement uncertainty due to the factors described.

A2.2	Definitions

Some terms are defined here for the purpose of this Annex.

Measurand

A quantity subjected to measurement.

Uncertainty

An estimate characterising the range of values within which the true value of a measurand lies.

Confidence level

The probability that the true value of the measurand lies within the range of values bounded by the uncertainty.

Design limits

The range of values within which a measurand should lie to ensure satisfactory operation.

A2.3	Effect of uncertainty upon the perceived design tolerance

Suppose that a design parameter has the value 10 dB with the design tolerances of +/- 2 dB and that the tester is perfect and has absolute precision, ie no uncertainty, then on test any indicated value between 8 dB and 12 dB will be acceptable; this can be said with complete confidence.

Suppose, however, that the tester is real and has an uncertainty of, say, +/- 1 dB when indicating any value within its range. At an indicated value of 8 dB, it can be stated that the true value of the measurand lies somewhere in the range 7 dB to 9 dB (ie 8 dB +/- 1 dB); similarly at an indicated value of 12 dB the true value of the measurand lies between 11 dB and 13 dB (i.e. 12 dB +/- 1 dB).

For an indicated value at the centre of the design range, in this example 10 dB, the true value of the measurand lies in the range 9 to 11 dB, which is entirely satisfactory. For some indicated values which lie off-centre the true value of the measurand still lies in the range 8 to 12 dB. For indicated values at the extremes of the design parameter, however, there is a significant probability that the true value of the measurand lies outside the design tolerances. There is, therefore, a finite probability that for an indicated value within design tolerance the true value actually lies outside the design parameter, e.g. in the lower sub-range 7 to 8 dB, or in the upper sub-range 12 to 13 dB.

Therefore for a pass result anywhere in the indicated range 8 dB to 12 dB the true value of the measurand lies between 7 dB and 13 dB. The indicated value will, however, be a single value not a range of values and for any single indicated value the uncertainty is +/- 1 dB.

2.4	Confidence level

The confidence level can be specified for a given magnitude of uncertainty 'u' expressed in terms of the standard deviation 's' of the distribution. The nature of the distribution of the uncertainties gives a clue as to this probability. For a Gaussian distribution with a standard deviation 's' the following, well known, standard results apply.

Estimated confidence probability	Uncertainty +/- u equivalent to

95.0 %	+/-1.96 s

95.5 %	+/-2.0 s

99.0 %	+/-2.58 s

99.7 %	+/-3.0 s

For example, where the estimated confidence probability is 95 % the true value of the measurand lies within a range of +/-1.96 s from the mean. Thus in the example above if the uncertainty +/- u is +/-1 dB, at the 95 % confidence level, then there is a 95 % probability that the true value of the measurand lies within +/-1 dB of the indicated value.

A2.5	Result of a test

From the foregoing it will be apparent that the result of a test can have several outcomes. The measurand may lie inside or outside the design limits. Where a result lies well within the design limits there is a high probability that the true value of the measurand also lies within the design limits. Where a result is exactly coincident with a design limit, there is still a significant probability that the true value of the measurand lies within the tolerance range but there is also a significant probability that the true value actually lies outside the range.

Where an indicated value lies outside the design limits but within the uncertainty range there is a high probability that the true value of the measurand also lies outside the design limits. However, there is a probability that the true value lies within the design limits. Depending upon the exact indicated value and its relation to the design limits, this probability can range from nearly 50 % for an indicated value just outside the design limits, to about 2.5 % for an indicated value at the edge of the design limits extended by the uncertainty; should this situation occur it should be noted.

Thus when a result lies close enough to a design limit for the distribution due to the measurement uncertainty to straddle that limit it is clear that the true value could lie on the other side of that limit. This poses a problem in interpretation and some rules have been formulated for the avoidance of doubt.

The result of a test shall be recorded in the test report in the following way:

1	Result within design limits.

2	Result outside the design limits.

If within the uncertainty this fact shall be stated.

Note:	Result 1 includes the case of exact coincidence with a limit.

A2.6	Receiver tests

Receiver tests present a particular difficulty due both to the statistical nature of the tests and to the rapid rate of change of performance with respect to stimulus level.

A very small change in stimulus level (e.g. within the setting error) can drastically affect the outcome of a test. Also the number of transmission errors due to the test propagation conditions must be statistically significant and at low error rates this impacts upon the duration of the test. These factors are catered for in the test method while table Ann. 2-1 shows the effect of the uncertainties of the applied stimulus level (column headed "Test limit with Uncert.") on the error rate. The values in this column are the estimated error rates at a stimulus level equal to level specified for the particular test minus the estimated uncertainty.

The outcome of each test shall be reported as described in Section 2.5 of this Annex.

GSM900 only - begin

Table Ann.2-1 : Effect of uncertainties on "Max-events"

Type of	Type of	Propagation/	Test limit	Test limit	Minimum No

test	channel	frequency	FER/BER %	FER/BER %	of

conditions		with uncert.	samples

BFI	TCH/FS	Static	0.244	same	82000

Sensitivity	TCH/FS	Static/FH	0.122	same	164000

''	TCH/FS Class Ib	Static/FH	0.410	same	20000000

''	TCH/FS Class II	Static/FH	2.439	same	8200

''	TCH/FS	TUhigh/NoFH	6.742	14.000	8900

''	TCH/FS Class Ib	TUhigh/NoFH	0.420	0.900	1000000

''	TCH/FS Class I	TUhigh/NoFH	8.333	9.000	120000

''	TCH/FS Class II	HT/NoFH	9.333	10.000	60000

''	TCH/FS Class II	RA/NoFH	7.500	8.000	24000

''	FACCH/F	TUhigh/NoFH	8.961	14.000	6696

''	FACCH/H	TUhigh/NoFH	7.728	12.000	7764

''	TCH/F9.6&H4.8	HT/No FH	0.778	2.000	180000

''	TCH/F4.8	HT/NoFH	0.011	0.040	5350000

''	TCH/F2.4	HT/NoFH	0.001	0.004	11900000

''	TCH/H2.4	HT/NoFH	0.011	0.040	5350000

Input Level	TCH/FS Class II	Static<-40dBm	0.012	same	1640000

range	TCH/FS Class II	Static<-15dBm	0.122	same	164000

TCH/FS Class II	EQ	3.250	same	120000

Co-channel	TCH/FS	TUlow/NoFH	24.000	25.100	25000

rejection	TCH/FS Class Ib	TUlow/NoFH	2.091	2.180	3300000

''	TCH/FS Class II	TUlow/NoFH	4.300	4.470	2000000

''	TCH/FS	TUhigh/FH	3.371	4.000	17800

''	TCH/FS Class Ib	TUhigh/FH	0.215	0.300	2000000

''	TCH/FS Class II	TUhigh/FH	8.333	9.000	1200000

''	FACCH/F	TUlow/NoFH	24.000	25.100	25000

''	FACCH/H	TUlow/NoFH	24.000	25.100	25000

''	TCH/F9.6 or H4.8	TUhigh/FH	0.336	0.424	178500

''	TCH/F4.8	TUhigh/FH	0.011	same	5350000

''	TCH/F2.4	TUhigh/FH	0.001	same	11900000

''	TCH/H2.4	TUhigh/FH	0.011	same	5350000

Adjacent	TCH/FS	TUhigh/NoFH	6.742	12.000	8900

channel	TCH/FS Class Ib	TUhigh/NoFH	0.420	0.800	1000000

200 kHz	TCH/FS Class II	TUhigh/NoFH	8.333	10.000	600000

''	FACCH/F	TUhigh/NoFH	10.640	16.900	5639

Adjacent	TCH/FS	TUhigh/NoFH	11.461	18.710	8900

channel	TCH/FS Class Ib	TUhigh/NoFH	0.756	1.670	1000000

400 kHz	TCH/FS Class II	TUhigh/NoFH	9.167	10.660	600000

''	FACCH/F	TUhigh/NoFH	19.152	30.500	3133

Intermod.	TCH/FS Class II	Static	2.439	same	8200

''	FACCH	TUhigh/NoFH	8.961	14.000	6696

Blocking &	TCH/FS Class II	Static	2.439	same	8200

spurious resp.	FACCH	TUhigh/NoFH	8.961	14.000	6696

GSM900 only - end

DCS1800 only - begin

Table Ann.2-1 : Effect of uncertainties on "Max-events"

Type of	Type of	Propagation/	Test limit	Test limit	Minimum No

test	channel	frequency	FER/BER %	FER/BER %	of

conditions		with uncert.	samples

BFI	TCH/FS	Static	0.244	same	82000

Sensitivity	TCH/FS	Static/FH	0.122	same	164000

''	TCH/FS Class Ib	Static/FH	0.410	same	20000000

''	TCH/FS Class II	Static/FH	2.439	same	8200

''	TCH/FS	TUhigh/NoFH	4.478	9.000	13400

''	TCH/FS Class Ib	TUhigh/NoFH	0.320	0.700	1500000

''	TCH/FS Class II	TUhigh/NoFH	8.333	9.000	120000

''	TCH/FS Class II	HT/NoFH	9.333	10.000	30000

''	TCH/FS Class II	RA/NoFH	7.500	8.000	24000

''	FACCH/F	TUhigh/NoFH	4.368	7.970	13736

''	FACCH/H	TUhigh/NoFH	7.752	10.800	7440

''	TCH/F9.6&H4.8HT	HT/NoFH	0.784	2.000	76500

''	TCH/F4.8	HT/NoFH	0.011	0.040	5350000

''	TCH/F2.4	HT/NoFH	0.001	0.004	11900000

Input Level	TCH/FS Class II	Static<-23dBm	0.122	same	164000

range	TCH/FS Class II	Static<-40dBm	0.012	same	1640000

TCH/FS ClassII	EQ	3.250	same	60000

Co-channel	TCH/FS	TUlow/NoFH	24.000	25.100	25000

rejection	TCH/FS Class Ib	TUlow/NoFH	2.091	2.180	3300000

''	TCH/FS Class II	TUlow/NoFH	4.300	4.470	2000000

''	TCH/FS	TUhigh/FH	3.371	4.000	17800

''	TCH/FS Class Ib	TUhigh/FH	0.215	0.300	2000000

''	TCH/FS Class II	TUhigh/FH	8.333	9.000	1200000

''	FACCH/F	TUlow/NoFH	24.000	25.100	25000

''	FACCH/H	TUlow/NoFH	24.000	25.100	25000

''	TCH/F9.6 or H4.8	TUhigh/FH	0.336	0.424	178500

''	TCH/F4.8	TUhigh/FH	0.011	same	5350000

''	TCH/F2.4	TUhigh/FH	0.001	same	11900000

''	TCH/H2.4	TUhigh/FH	0.011	same	5350000

Adjacent	TCH/FS	TUhigh/NoFH	3.371	6.740	17800

channel	TCH/FS Class Ib	TUhigh/NoFH	0.270	0.540	2000000

200 kHz	TCH/FS Class II	TUhigh/NoFH	8.333	9.580	1200000

''	FACCH/F	TUhigh/NoFH	3.808	6.950	15756

Adjacent	TCH/FS	TUhigh/NoFH	5.714	10.140	10500

channel	TCH/FS Class Ib	TUhigh/NoFH	0.448	0.780	1200000

400 kHz	TCH/FS Class II	TUhigh/NoFH	9.167	10.760	720000

''	FACCH/F	TUhigh/NoFH	6.832	13.160	8782

Intermod.	TCH/FS Class II	Static	2.439	same	8200

''	FACCH	TUhigh/NoFH	4.368	7.970	13736

Blocking &	TCH/FS Class II	Static	2.439	same	8200

spurious resp.	FACCH	TUhigh/NoFH	4.368	7.970	13736

DCS1800 only - end

A2.7	Pass / fail criteria

The outcomes of these tests might be interpreted in terms of pass / fail but it is not the function of the system simulator (SS) to do this interpretation. The SS reports the results of the tests.